## Slide 1: SE108 Support information
1/20/15

## Slide 2: SE108 Enclosure specs
Depth rating:70 m (100 psi)
Dimensions:
  Weight: ~20 kg (~45 lbs)?
  Length: 0.7 m (27.5”)?
  Diameter: 24 cm (8.5”)
Materials:
  Housing: White PVC sch. 80
  Endcaps: Black Delrin
  Subconn circular micro connectors (brass, neoprene)

## Slide 3: SE108 hardware specs
Electrical (See CR1000 manual for datalogger specifics):
  Voltage range
    Input: 9-16 VDC
    Output: 11-16 VDC, unregulated
  Power consumption
    Sampling: depends on sensor load and configuration (sleep + 30mA / sensor)?
    Sleep:20mA?
  Current limits
    Input:10A (ATM fast acting) - Power in
      One power input port – Teledyne impulse VSG-2-BCL
    Output: unregulated voltage at 1-5A depending on port (*user selectable/field serviceable)
      Seven sensor ports
Sensor port layout and configuration
  1 host/PC RS232 port – Port 1
  7 sensor ports
    EXO (DCP-SOA sdi-12)- Port 2
    Analog (1 SE channel, 0-5VDC)- Port 6
    Three serial (2x sc105, 1x control ports 3 and 4)- Ports 3,5,8
    One sdi12 port - Port 7
    One ZT wiper* (90 degrees unless otherwise specified)-Port 4
*Power fused and switched to all sensor ports by SDM-CS8s and SSR

## Slide 4: SE108 bulkhead port assignments
COM3 serial port
Serial host
SC105 serial port (SDC8)
Carrying handle
ZT Wiper port
EXO DCP-SOA
SDI12 port
SC105 serial port (SDC7)
Analog port
Main Power port

## Slide 5: SE108 internals (power side)
Main power input
Fuse array
SC105 serial interface (2x)
SDM relay board
Solid state relay array

## Slide 6: SE108 internals (control side)
NL115
Wiper driver board
CR1000
DCP-SOA
CF card slot

## Slide 7: Wiring diagram